# SOP for making an ALD run – in its simplest form:

- Make sure that you are authorized to use this instrument. If you are a bit unsure – contact the instrument responsible.

- Read the manual handed to you during the training.

- Make sure that the reactor is at atmospheric pressure.

- Load the substrates and the precursors.

  - If the solid precursor is toxic/pyrophoric fill it into the boat in a fume hood or in a glove box and use a sealed transport tube for transferring the boat into the reactor.

  - If you have pyrophoric/toxic precursors in a bubbler make sure that all the connections in the line are tight and made properly and leak test the line before opening the main valve.

- Close the reactor and click pump down.

- Load the relevant run file and start heating

- Open the precursor main valves manually.

[WAIT FOR THE RUN TO FINISH]

- Close the manual valves of the bubblers when the run is done.

- Pump up the reactor when the temperature is sufficiently low (depending on your precursors)

- Open the reactor and take out the substrates.

- Close the reactor.

- If you are using toxic chemicals make sure that the reactor, including the pump line is perfectly vacuum tight.

- This SOP is for a simple ALD-run and other details may apply to your specific process/experiment and the procedure must therefore be adapted in discussion with instrument responsible and supervisors.

# Risk assessment for Beneq TFS500

| Subtask | Risk | Precaution |
| --- | --- | --- |
| Loading in substrates and precursors | The reactor is still hot/heating or under vacuum | Use your head, and heat proof gloves. |
| Loading in a solid precursors in a boat | You get exposed to toxic precursors during precursor loading | Use a bubbler or fill the precursor into the boat in a fume hood or in a glove box and transfer it to the reactor using a sealed transport tube. |
| Attaching a pyrophoric/toxic precursor in a bubbler (This is not standard procedure for normal user) | The chemical leaks from the bubbler to the atmosphere. | Make sure that you connect the bubbler properly, that the all VCR connections are made correctly and that you use proper gaskets and sealing that are compatible with the chemical. Leak test the line if needed. Do not open any valves before the bubbler is properly connected and the reactor is under vacuum. |
| Using a toxic precursor | The precursor may leak to the atmosphere from the exhaust | Use a gas scrubber for the specific chemical. |
| Write a run file and start the deposition. | Incompatible chemicals may cause fire or other dangerous reactions in the reactor or in the pump. | Make sure that you don’t mix chemicals that form explosive mixtures (for example H2 and O2) or that react with the pump oil (if you use oil pump). Do your job with respect to risk assessments |
| Running a deposition | Toxic/harmful chemicals can leak to the atmosphere. | If you are using toxic precursors make sure that the reactor, including the exhaust line, is not leaking to the atmosphere and that the ventilation in the pump room works properly. |
| The run is over | The valves to the bubblers are leaking and the reactor is filled with the chemicals. | Close the manual valves to the bubblers when the run is over. |
| Removing bubbler containing toxic/pyrophoric precursors. | The precursors leaks to the atmosphere. | Close the manual valve and evacuate the tubing down to the manual valve before disconnecting the bubbler. |

# Also, check out the document “Examples, growth rate vs temperature Al2O3, ZnO, TiO2” to compare test runs.